## Sheet1
PPM Gas Testing | [6 empty] | Procedure:
| Unit 1 |  | Unit 2 |  | Unit 3 |  | 1. Qualify the units
PPM Concentration | Before -> After Calibration | Measure | Before -> After Calibration | Measure | Before -> After Calibration | Measure | 2. Clear calibrations
First Test | [6 empty] | 3. Calibrate and record reading on an Airgas 8ppm tank (green)
7.95 PPM (Airgas/Control) | 10.51 PPM -> 7.95 PPM |  | 13.50 PPM -> Fail (Unable to span) |  | 8.50 PPM -> 7.95 PPM |  | 4. Switch to Zero gas
8 PPM (West Air) |  | 8.1 PPM |  |  |  | 8.2 PPM | 5. Measure West Air (Silver) 8ppm, record reading
76.5 PPM (West Air) |  | 74.2 PPM |  |  |  | 78.1 PPM | 6. Switch to Zero gas
7. Measure West Air (silver) 80ppm, record reading
Second Test | [6 empty] | 8. Switch to Zero gas
81.3 PPM (Airgas/Control) | 91 PPM -> 81.3 PPM |  | 118.8 PPM -> 81.3 PPM |  | 77.5 PPM -> 81.3 PPM |  | 9. Calibrate and record reading on Airgas 80ppm tank (green)
8 PPM (West Air) |  | 9.08 PPM |  | 8.68 PPM |  | 8.74 PPM | 10. Repeat steps 4-8.
76.5 PPM (West Air) |  | 86.5 PPM |  | 86.5 PPM |  | 86.5 PPM
Other procedure done:
Third Test | [6 empty] | Calibrate and record reading on West Air 80 ppm tank (silver)
76.5 PPM (West Air/Control) | 97 PPM -> 76.5 PPM |  | 126.9 PPM -> Fail (Unable to span) |  | 82.7 PPM -> 76.5 PPM |  | Measure Airgas 80 ppm, record reading
81.3 PPM (Airgas) |  | 72 PPM |  |  |  | 71.9 PPM | Measure West Air 8 ppm, record reading
8 PPM (West Air) |  | 8.02 PPM |  |  |  | 7.77 PPM
Note: The test was performed on GPR-1200 units with GPR 12-333 sensors.
Findings: From the test that I performed to check how accurate the new gas from the vendor West Air was, I have concluded that the gas appears to be more accurate on both levels, the 8 PPM and 76.5 PPM concentrations. The 8 PPM gas was very constant; the most division obtained is no more than 2 PPM. The 76.5 ppm was also very constant, with the most divitation being less than 5 PPM. Also finding that the 81.3 PPM gas from Airgas might not be correct with me getting a reading of approximately 10 PPM less when checked the same way as the other gases.